## Slide 1: dVRK PI / User Group Meeting
September 26, 2023
(online)

## Slide 2: Agenda
Introduction
Announcements
Community Presentation: Allison Okamura (Stanford Univ.)
dVRK/dVRK-Si Hardware/Software Updates
Controllers (dVRK and dVRK-Si)
FPGA V3
Firmware updates
Embedded software
PC software
Open discussion

## Slide 3: Announcements
Augmented Dexterity: Ken Goldberg (UC Berkeley)
2023-2024 AccelNet Surgical Robotics Challenge

## Slide 4: AccelNet Surgical Robotics Challenge
Relaunch of 2021-2022 challenge
Online phase
  AMBF Simulator
In-person phase (JHU)
  Summer 2024?
Focus on surgical autonomy:
  Needle detection
  Needle driving
  Suturing
This video created via teleoperation
Challenge: Can you do it autonomously?
OISE-1927275 (WPI)
OISE-1927354 (JHU)
https://surgical-robotics-ai.github.io/surgical-robotics-challenge-2023/challenge-2023.html

## Slide 5: Community Presentation
Allison Okamura, Stanford University

## Slide 6: Agenda
Introduction
Announcements
Community Presentation: Allison Okamura (Stanford Univ.)
dVRK/dVRK-Si Hardware/Software Updates
Controllers (dVRK and dVRK-Si)
FPGA V3
Firmware updates
Embedded software
PC software
Open discussion

## Slide 7: dVRK and dVRK-Si Controllers
dVRK
(V1-V3)
dVRK-Si
da Vinci S
da Vinci “Classic”
da Vinci Si
(some)

## Slide 8: dVRK Controller (V1 & V2)
E-Stop
dVRK Controller
QLA
QLA
dMIB
FPGA
FPGA
Legend
Ethernet pass-through
Ethernet cable
FireWire pass-through
FireWire cable
E-Stop connector
E-Stop wiring
Data cable (SCSI 68)
Power cable (DB 9)
Logic Power Supply
Motor Power Supplies
MTM (12V + 24V) PSM(24V) ECM (36V)
Front LEDs
Arm
FireWire
Ethernet
V1: FireWire only
V2: FireWire + Ethernet
(could also use FPGA V3)

## Slide 9: dVRK Controller (V3)
dVRK Controller
QLA
QLA
DQLA-F
dMIB
FPGA V3
Legend
Flat flex cable
Data cable (SCSI 68)
Power cable (DB 9)
Notes
Ethernet and FireWire external connections use PCB-mounted connectors on FPGA
E-Stop external connections use PCB-mounted connector on DQLA-F
Logic Power Supply
Motor Power Supplies
MTM (12V + 24V) PSM(24V) ECM (36V)
DQLA-Q
Front LEDs
E-Stop
Arm
FireWire
Ethernet

## Slide 10: dVRK-Si Controller
Si Controller
dRAC
FPGA
Logic Power Supply
Motor Power
Supply
(48V)
Front LEDs
Notes
Ethernet and FireWire external connections use PCB-mounted connectors on FPGA
Cannot use FPGA V1 or V2 (not enough I/O)
E-Stop external connections use PCB-mounted connectors on dRAC
Power and serial communication (LVDS) to the arm use PCB-mounted connectors on dRAC
E-Stop
Arm
Ethernet
FireWire

## Slide 11: dVRK-Si Controllers
Standalone arms: Controller + Cables + Arm + ESPM Programmer
Full system with SUJ: Controller + dSIB-Si + dESSJ + Patient Cart + ESPM Programmer

## Slide 12: dVRK-Si Setups

## Slide 13: ESPM Programmer
Reprograms the ESPM (Altera FPGA).
Sticks to the arm. Connects to the JTAG port of ESPM.
Replays an XSVF file stored in the SD card when the arm powers up.
Firmware update by replacing the file in the SD card.
Reprogramming is not persistent. Unplugging or toggling switch, and then power cycling, restores original firmware.

## Slide 14: Controller Build Status
dVRK Controllers: Update from Greg Fischer
dVRK-Si Controllers: 12 controllers available

## Slide 15: Agenda
Introduction
Announcements
Community Presentation: Allison Okamura (Stanford Univ.)
dVRK/dVRK-Si Hardware/Software Updates
Controllers (dVRK and dVRK-Si)
FPGA V3
Firmware updates
Embedded software
PC software
Open discussion

## Slide 16: FPGA V3.1
Upgrade from Spartan 6 to Zynq (Spartan 7)
Bigger FPGA + dual-core ARM processor
FPGA
Key features:
Backward compatible with prior FPGA versions
Additional Ethernet port (2 total, 1GB each)
Linux on ARM processor (1 GB DDR3 RAM)
Micro-SD for firmware upgrades
4 additional I/O (V3.1)
Status:
Built 108 boards
10 failed permanently (9% failure rate)
13 with minor issues
DRAM
Ethernet x 2
Micro-SD
Firewire x 2
FPGA

## Slide 17: Network Connections
FPGA
FPGA
FPGA
FPGA
PC
real-time connection
(Firewire or Ethernet/Firewire)
non-real-time connection
(Ethernet)
Firewire
Ethernet
Note: Ethernet daisy-chain not yet supported by firmware
EtherCAT could be an option (would need to license IP core)

## Slide 18: Agenda
Introduction
Announcements
Community Presentation: Allison Okamura (Stanford Univ.)
dVRK/dVRK-Si Hardware/Software Updates
Controllers (dVRK and dVRK-Si)
FPGA V3
Firmware updates
Embedded software
PC software
Open discussion

## Slide 19: Firmware Rev 8
Released August 10, 2023
Required for FPGA V3
For FPGA V1 and V2:
  Required to co-exist with FPGA V3, including dVRK-Si
  Required to use QLA V1.6
  Supports “dongle” for reading instrument id
    Required for controller builds 1-4 (2012-2015)
    Distributed at ICRA and Hamlyn 2022
https://github.com/jhu-dvrk/sawIntuitiveResearchKit/wiki/Tool-Detection

## Slide 20: Firmware V8 Configurations
 | QLA (dVRK, V1/V2) | DQLA (dVRK, V3) | dRAC (dVRK-Si)
FPGA V1 | FPGA1394V1-QLA | Not feasible | Not supported
FPGA V2 | FPGA1394V2-QLA | Not feasible | Not supported
FPGA V3 | FPGA1394V3-QLA | FPGA1394V3-DQLA | FPGA1394V3-DRAC
Notes:
FPGA1394V1 previously called FPGA1394 (Rev 1-7)
FPGA1394V2 previously called FPGA1394Eth (Rev 1-7)
DQLA not feasible with V1/V2 due to lack of 4 extra I/O introduced with V3
dRAC not officially supported with V1/V2, but is expected to work for arm only (SUJ requires 4 extra I/O introduced with V3)

## Slide 21: Firmware Update
For FPGA V1 and V2, continue to flash EPROM via FireWire (or Ethernet) using pgm1394
  https://github.com/jhu-cisst/mechatronics-firmware/wiki/FPGA-Program
For FPGA V3, copy release to MicroSD card
  https://github.com/jhu-cisst/mechatronics-embedded/releases
  Utility: dvrk-sd-card-updater.py
  Documentation forthcoming

## Slide 22: Firmware V9 Goals
Fix Ethernet support for FPGA V3
Migrate from Xilinx ISE to Vivado for FPGA V3
  FPGA V1 and V2 only supported by ISE
Support digital current control for QLA V1.5+
Add digital (FIR) filters on analog feedback

## Slide 23: Agenda
Introduction
Announcements
Community Presentation: Allison Okamura (Stanford Univ.)
dVRK/dVRK-Si Hardware/Software Updates
Controllers (dVRK and dVRK-Si)
FPGA V3
Firmware updates
Embedded software
PC software
Open discussion

## Slide 24: Embedded Software
FPGA V3 has Xilinx Zynq System on Chip (SoC), which has FPGA (PL) and dual-core ARM processor (PS)
PS
PL
Zynq
2 x eth
2 x fw
Companion Board
(QLA, DQLA, dRAC)
Embedded software
Firmware

## Slide 25: Embedded Software
GitHub: https://github.com/jhu-cisst/mechatronics-embedded
Dual-core ARM processor, running embedded Linux (6.1.5 kernel)
  Built using Xilinx Petalinux tool (V2023.1)
  Custom library (libfpgav3) and applications (fpgav3init, fpgav3sn)
Initialization process (fpgav3init)
  Detect companion board type, then load correct firmware
  Read FPGA serial number
  Program flash (if needed)
  Other initialization

## Slide 26: Embedded Software – Initialization
 | QLA | DQLA | dRAC
FPGA V3 | FPGA1394V3-QLA | FPGA1394V3-DQLA | FPGA1394V3-DRAC
Also: FPGA1394V3-BCFG (boot configuration firmware to determine board type)
Power Up
Load
BCFG
Determine board type
Load QLA
Load DQLA
Load DRAC
Boot Linux
Will keep BCFG firmware if QLA, DQLA or DRAC not detected
- Most likely cause: MicroSD card not inserted
Can still talk to board via FireWire or Ethernet, but will not drive outputs
Allows the same MicroSD card to be used everywhere!

## Slide 27: Embedded Software – Utilization
Can cross-compile applications to run on embedded Linux
  Advantage: high-speed interface to FPGA logic
  Disadvantage: lower performance and less memory than PC
Current limitation: FPGA processor not easily accessed
  Ethernet connection to processor not yet supported
  6-pin serial console connector not exposed (inside controller box)

## Slide 28: Embedded Software – Cross-Compiling
Option 1: Use toolchain_vitis_fpgav3.cmake, created in mechatronics-embedded build directory
Option 2: Use Vitis IDE
When configuring mechatronics-embedded (option 1) or Vitis IDE (option 2), choose sysroot as:
  Released sysroot (on GitHub)
  Locally built sysroot (using Petalinux)
Documentation forthcoming
For an example of option 1, see mechatronics-software repository (low-level PC software)

## Slide 29: Agenda
Introduction
Announcements
Community Presentation: Allison Okamura (Stanford Univ.)
dVRK/dVRK-Si Hardware/Software Updates
Controllers (dVRK and dVRK-Si)
FPGA V3
Firmware updates
Embedded software
PC software
Open discussion

## Slide 30: dVRK Software Updates
Improvements
New hardware (dVRK-Si + DQLA)
ROS
Release

## Slide 31: Software: better control
PID
  IO and PID components now use actuator space
  Added desired velocity, better trajectory following
  Added disturbance observer, better positioning
  Fixed deadband implementation
Teleoperation
  Updated to use the MTM cartesian position and velocity to send a position and velocity goal to the PSM
  Promising results: “latency” between MTM and PSM down from
    ~100ms to ~60ms with better PID gains
    ~60ms to ~10ms with Cartesian velocity

## Slide 32: Software: miscellaneous
Rewrote configuration generator
  Support new hardware (dRAC, DQLA)
  Python based, no need to install Matlab
  Generates IO, arm and console configuration files
SUJs
  Cleaned up code for Classic SUJ
  Added support for Si SUJ
  Added “fixed” SUJ class to support groups with PSMs and ECM on custom frames, related poses can come from hand-eye calibration
More Python examples

## Slide 33: Software: support for new hardware
Si arms (PSM and ECM)
  Brakes on PSM Si
  Different code to get instrument model
  New potentiometer calibration procedure
  Different configuration files for IO, PID and kinematics
  New configuration file generator, Python based
  Modified console class to accommodate generations, Classic vs. Si
Si SUJ
  Bluetooth used to get all SUJ joint values
  SUJ brakes are working
  Forward kinematics implemented
  No good method for potentiometer calibration

## Slide 34: Software: ROS 1 and ROS 2
Past years updates
  dVRK software built and tested on Ubuntu 20.04 with ROS 2 Galactic and 22.04 with ROS 2 Humble
  Missing launch files, calibration scripts, CRTK Matlab client… so ROS 1 was needed to get started
ROS 2 without ROS 1
  Ubuntu 20.04 is the last version with ROS 1 support and is getting a bit old (though EOL is still 04/25)
  Upcoming dVRK release to support ROS 2 without any need for ROS 1 so one can use Ubuntu 22.04 with ROS 2 Humble or later
But we still need to maintain ROS 1 so lots of code duplication…

## Slide 35: Software: ROS 1 and ROS 2
Server side:
  cisst to ROS 1 bridge ported to ROS 2
  Code can’t be easily shared but we can manually maintain both code bases in synch using diff tools
Client side using CRTK clients
  Extended CRTK Python and Matlab client libraries to introduce a ROS Abstraction Layer (RAL) so scripts can be ROS version agnostic
  Updated all existing scripts (calibration, utilities and examples) to use crtk.ral
Status
  Python client library updated and fairly well tested
  Matlab client library updated but need further testing
  CRTK documentation partially ported to readthedocs: https://crtk-robotics.readthedocs.io/en/latest/

## Slide 36: Software: release
Aiming for November!
To do:
Finish porting a few Python scripts
Test Matlab CRTK for MTM gravity compensation
SUJ Si calibration method
Create new ROS packages for video pipeline
Documentation! None of the new features for the Si are documented  Likely to use readthedocs
For groups receiving hardware before upcoming release:
Development branches are fairly stable
Anton will help over Zoom

## Slide 37: Open Discussion